| Pilot NMI logo | NMI descriptor etc. |
| --- | --- |

<Executive Report Template v2 – 25/8/17 – Delete this line>

CCL Key Comparison

Measurement of gauge blocks by interferometry

<local RMO identifier, as required>

CCL-K1

Executive Report

A Pilot (NMI),

Location, Month Year

Contents

<Insert automatic table of contents, using heading style Heading 1 to Heading 2, as appropriate>

# Document control

Version 2, issued on 12 Nov 1999.

# Introduction

At its meeting in September 1997, the Consultative Committee for Length, CCL, decided upon a key comparison on gauge block measurements by interferometry, named CCL-K1, with METAS as the pilot laboratory. The comparison final report was approved in May 1999. This is the Executive Report on the comparison.

# Scope of the comparison

Ten gauge blocks made from steel and ten made from tungsten carbide were circulated in a single loop. The artefacts ranged in size from 0.5 mm to 100 mm. Artefact circulation started in March 1998 and was completed in March 1999.

# Participants

Give a concise list of the participants.

Table 1. List of participant laboratories and their contacts.

| Laboratory Code | Laboratory name | Contact & email |
| --- | --- | --- |
| NPL | National Physical Laboratory | Andrew Lewis Andrew.Lewis@npl.co.uk |
| LNE | Laboratoire National d’Essais | Georges Vailleau GVailleau@lne.fr |
| NIST | National Institute of Standards and Technology | John Stoup John.Stoup@nist.gov |

## Changes in participation

Clearly state any participants which joined the comparison later and those that withdrew from the comparison at any stage, listing the reasons for withdrawal.

# Processing of the results

- The results were analysed using the weighted mean as the Key Comparison Reference Value.

- Some results were excluded from the calculation of the weighted mean as they were determined to be outliers. Determination of outliers was initially performed through statistical analysis. There was then discussion to determine if the laboratory concerned could explain the problem. By such a technique, four results from one participant and one from another participant were identified as outliers and removed from contributing to the weighted means.

- One participant withdrew all of their results due to equipment failure.

# Implications regarding the MRA and CMC claims

There should be a listing of the participants’ results and the uncertainties in the format of degrees of equivalence, DoE, and the uncertainty of each DoE. This should be accompanied by a listing of any CMCs which are already in the KCDB which relate to the subject of the comparison. This should then be followed by a discussion of the results and their impact on any existing or future CMC claims. Follow the format and instructions given in CCL-WG-MRA/GD3.

Different pilots present the data in different ways, as appropriate for the comparison. Here are some examples. Other pilots have used graphs where the measurands are complex.

Table 2 Differences of measured diameters with respect to the KCRVs and with the associated expanded uncertainties (k=2). Differences and uncertainties are in nm, the uncertainty is rounded up to the nearest nm. Figures in blue in braces indicate the CMC claim ( in the KCDB) of the NMI at the time of preparation of this Executive Report. [N – indicates no CMC currently listed in the KCDB]. Grey shaded numbers indicate differences from the KCRV by more than the k =2 uncertainty.

Recommendation 1

Unless mentioned otherwise below, the results of the participants are well controlled and support their existing CMC claims.

Recommendation 2

From Table 5, where the laboratory has a tick, all results of that parameter across the artefacts were accepted and the laboratory can claim a new CMC. It is left to any CMC submission for the laboratory to have a range and uncertainty peer reviewed. The last column of Table 5 lists the number of different parameters that the laboratory can claim. Therefore, all of the sixteen participants would be expected to be able to submit CMC claims except for SCL who only submitted results for parameters with an existing CMC listing.

Looking at one laboratory, NMIA had 1 listed CMC but of fifteen different parameters, had one result excluded. This suggests that NMIA could justify thirteen new CMC claims where the range of parameter and best uncertainty would have to be considered on individual results, noting that some parameters were reported on multiple artefacts for different sizes.

Recommendation 3

Six of the outliers are from NPLI where their submitted uncertainty, adjusted to k = 2, is less than the KCDB quoted CMC uncertainty (in each case, 5% of the parameter). If the CMC uncertainty is used in the relevant KCRV assessment, all six of these NPLI results would still have been removed from KCRV calculations. This suggests that NPLI should investigate both results and uncertainties for certain parameters and what impact these results have on the claims in the KCDB.

Recommendation 4

An adjusted En ratio does not guarantee that the result is excluded as the Birge ratio governs the limits of the population. This is the case for the KRISS outlier, Ra, shown in Table 6. If the CMC listed uncertainty was used, the new En = 1.07 but this result would remain in the population since calculations show the Birge ratio is within the Birge Criterion. Therefore, for this parameter, KRISS could justify maintaining this CMC with the CMC listed uncertainty.

Recommendation 5

The NIST outlier is a unique case where a stylus radius correction is applied which significantly increases the Ra value as discussed on page 8 and in more detail in the full report. If this correction was not applied, as no other laboratory does, the reported value of 0.9497 μm would be 0.9377 μm and using the submitted uncertainty of ± 0.0088 μm; this result would give an En of 0.41 and would not be excluded from the KCRV. Therefore NIST could therefore actually reduce their CMC uncertainty if backed up by a revised uncertainty budget supporting the new claim.

Recommendation 6

Where Table 5 shows a solid green cell with no tick or number, the laboratory results were accepted and support the existing CMC listing on the KCDB. Differences between submitted and listed uncertainties are not examined for these parameters. The laboratory may choose to leave the listing as it is or if a smaller uncertainty was used, may submit a reduced uncertainty if backed up by a revised uncertainty budget supporting the new claim.