RECOMMENDATION  ITU-R  TF.1153-3

The operational use of two-way satellite time and
frequency transfer employing PN codes

(Question ITU-R 201/7)

(1995-1997-2003-2009)

The ITU Radiocommunication Assembly,

considering

a)	the demonstrated high-accuracy for time and frequency comparisons using the two-way satellite time and frequency transfer (TWSTFT) method as expressed in Question ITU-R 201/7;

b)	the well established use of TWSTFT systems in global networks using telecommunication satellites in the Ku band [10.7 – 14.5 GHz], in support of, but not limited to, the realization of Coordinated Universal Time UTC;

c)	that other frequency bands are becoming important;

d)	that some TWSTFT links have been repeatedly calibrated so that time transfer with a systematic uncertainty below 1 ns could be achieved;

e)	that theoretical background is available to calculate the corrections for the effect of the propagation delay through the troposphere and the ionosphere, the correction for the Sagnac-effect, and other reciprocity factors;

f)	that TWSTFT is nowadays performed in networks with more than ten participating stations using the code division multiplex access;

g)	the need for standardizing:

–	measuring procedures,

–	data processing,

–	format for the exchange of data and relevant information between participating stations and interested bodies, such as the International Bureau for Weights and Measures (BIPM) ;

recommends

1	that the measuring and data processing procedures for accurate time and frequency transfer via TWSTFT be followed as outlined in Annex 1;

2	that the data formats for the exchange of the relevant data between participating stations and interested bodies be as outlined in Annex 2 .

Annex 1

Procedures for TWSTFT

# 1	Introduction

Radio links can be used to transfer time from one clock to another. In radio links however, the signal delays are changing with distance, ionosphere, troposphere, temperature, earth conductivity and so on. To cancel these influences to first order the two-way scheme has been introduced: at both clock sites the time signals are transmitted at nominally the same instant and on both sides the signal from the other clock is received and its time of arrival is measured. After the exchange of the measured data, the difference of the two clocks is calculated. The delays cancel due to the complete reciprocity, to first order, of the signal paths. The accuracy of the result then depends on the residual effects due to the incomplete reciprocity. Some of these effects are known and others are still unknown. In some cases corrections for these effects can be used to improve accuracy.

2. Brief description of the exchanged signals

TWSTFT is based on the exchange of timing signals through geo-stationary telecommunication satellites as it is schematically shown in Fig. 1 below. It is done by transmission and reception of radio frequency (RF) signals containing the pseudorandom noise (PRN) codes binary phase-shift keying (BPSK) modulated on the intermediate frequency (IF) by a modem. The phase modulation is synchronized with the local clock, and the modem generates a 1PPS output, synchronous with the BPSK sequence, and named 1PPSTX. This signal is the realization of a time scale named TS(k). Each station uses a dedicated PRN code for its BPSK sequence in the transmitted signal. The receiving equipment generates the BPSK sequence of the remote stations and to reconstitute a 1PPS tick from the received signal, named 1PPSRX. The difference TI(k) between the two 1PPS signals is measured by a time-interval counter (TIC). Following a pre-arranged schedule, a pair of stations lock on the code of the corresponding remote station for a specified period, called a session, measure the signal’s time of arrival, and store the results. After exchanging the data records the difference between the two clocks can be computed. In the remaining of this Annex 1, details of the data reduction and the treatment of systematic effects involved are elaborated.

Fig. 1 illustrates the various signal delays and explains how the time difference between the clocks at stations 1 and 2 can be determined.

TS(k):		Local time scale, physically represented by the 1PPSTX signal generated by the modem, k being 1 for station 1 and 2 for station 2

TI(k):		Time interval reading, the counter gate is opened by a 1PPS signal related to the local transmit signal and closed by a 1PPS signal related to the received signal; short-form designation of 1PPSTX – 1PPSRX

TX(k):		Transmitter delay, including the modem delay

RX(k):		Receiver delay, including the modem delay

SPU(k):		Signal path uplink delay

SPD(k):		Signal path downlink delay

SPT(k):		Satellite path delay through the transponder

SCU(k):	Sagnac correction in the uplink

SCD(k):	Sagnac correction in the downlink.

The difference of the time scale at station 2 from the time scale at station 1 expressed by TS(1) _ TS(2) is determined as follows.

The TIC reading at station 1 is:

TI(1)  TS(1) – TS(2)  TX(2)  SPU(2)  SCU(2)  SPT(2)  SPD(1)  SCD(1)  RX(1)

The TIC reading at station 2 is:

TI(2)  TS(2) – TS(1)  TX(1)  SPU(1)  SCU(1)  SPT(1)  SPD(2)  SCD(2)  RX(2)

Subtracting the expression for station 2 from that for station 1, gives:

TI(1) – TI(2)  2TS(1) – 2TS(2)  TX(2) – TX(1)  SPU(2) – SPU(1)  SPT(2) – SPT(1)
     SPD(1) – SPD(2)  RX(1) – RX(2)  SCD(1) – SCU(1) – SCD(2)  SCU(2)

The time-scale difference is thus given by:

TS(1) – TS(2)  1/2[TI(1)]	(	TIC reading at station 1)

1/2[TI(2)]	(	TIC reading at station 2)

1/2[SPT(1) – SPT(2)]	(	Satellite delay difference)

1/2[SCD(1)  SCU(1)]	(	Sagnac correction for station 1)

1/2[SCD(2)  SCU(2)]	(	Sagnac correction for station 2).

1/2[SPU(1)  SPD(1)]	(	Up/down difference at station 1)

–1/2[SPU(2)  SPD(2)]	(	Up/down difference at station 2)

1/2[TX(1)  RX(1)]	(	Transmit/receive difference at station 1)

–1/2[TX(2)  RX(2)]	(	Transmit/receive difference at station 2)

The last seven terms are the corrections for non-reciprocity. Except of the satellite delay difference SPT, the corrections can in principle be grouped in corrections per station.

The non-reciprocity factors are further addressed in the following sections.

# 3	Causes of non-reciprocity and systematic uncertainty in TWSTFT

# 3.1	Non-reciprocity due to satellite equipment delays

When the satellite receive antenna, transponder channel and transmit antenna are common to both signal paths, the satellite signal delays are equal, i.e. SPT(1)  SPT(2). This is not the case when different frequencies, transponders or different spot beams are used for the reception and/or transmissions from each station, e.g. using the transatlantic Intelsat satellites. In this case SPT(1) and SPT(2) or at least the difference SPT(1)  SPT(2) should be measured before the launch of the satellite or using another accurate method.

# 3.2	Sagnac-effect correction

Due to the movement of both the earth stations and the satellite around the rotation axis of the Earth during the propagation of a time signal to and from the satellite, a correction has to be applied to the propagation time of the signal. The Sagnac correction for the one-way path from satellite s to ground station k is given by:

SCD(k)  ( / c2)  [Y(k)  X(s) – X(k)  Y(s)]

where:

 :	Earth rotation rate  7.2921  10–5 rad/s

c :	speed of light  299 792 458 m/s

X(k) :	geocentric x-coordinate of station (m)

	r cos[LA(k)]  cos[LO(k)]

X(s) :	geocentric x-coordinate of satellite (m)

	R cos[LA(s)]  cos[LO(s)]

Y(k) :	geocentric y-coordinate of station (m)

	r cos[LA(k)]  sin[LO(k)]

Y(s) :	geocentric y-coordinate of satellite (m)

	R cos[LA(s)]  sin[LO(s)]

r :	Earth radius  6 378 140 m

R :	satellite orbit radius  42 164 000 m

LA :	latitude

LO :	longitude.

For geo-stationary satellites LA(s)  0, so:

SCD(k)  ( / c2)  R  r  cos[LA(k)]  sin[LO(k) – LO(s)].

The total Sagnac correction SCT(1,2) for a measurement of the clock at station 2 with reference to the clock in station 1 is:

SCT(1,2)  ½[SCU(1)  SCD(2) – [SCU(2)  SCD(1)]]

Furthermore, the sign of the Sagnac correction for the downlink is opposite to the sign of the Sagnac correction for the uplink due to the opposite directions of the signals: SCU(k)  –SCD(k), so that SCT(1,2)  –SCD(1)  SCD(2) is valid.

Example for a satellite at 307 E:

LA(VSL)  52 N, LO(VSL)  4 E, LO(sat)  307 E, difference in LO  57, SCD(VSL)  112.42 ns

LA(USNO)  39 N, LO(USNO)  283 E, LO(sat)  307 E, difference in LO  –24, SCD(USNO)  – 68.83 ns

SCT(VSLUSNO): –SCD(VSL)  SCD(USNO)  –181.25 ns

SCT(USNOVSL): –SCD(USNO)  SCD(VSL)  181.25 ns

VSL:	NMi Van Swinden Laboratory, Delft, the Netherlands

USNO:	US Naval Observatory, Washington DC, United States of America.

The position of a geo-stationary satellite is not perfectly fixed with respect to the observer on earth. A small periodic motion with a daily period occurs around a central position. This causes a periodic variation of the Sagnac effect with a maximum peak to peak amplitude of a few hundred of ps. This has been mostly ignored up to now, but needs to be accounted for if particularly high accuracy is aimed at.

## 3.3	Path delay difference due to satellite movement in an earth fixed frame

Two-way paths between earth stations via the satellite are not reciprocal if the satellite is in motion relative to the Earth’s surface and if the two arriving signals do not pass through the satellite at the same instant. This effect can in principle be avoided by an offset between  the emission times at the two stations to compensate for the different path delays SPU(1) and SPU(2) to the satellite. If the signals from the two stations arrive at the satellite within 5 ms, the delay difference is at few tens of ps level, and it shows a diurnal pattern.

## 3.4	Ionospheric correction

The up and downlink signals at each station differ in carrier frequency and they experience a different ionospheric delay equal to:

40.3 TEC (1c)(1fd2−1fu2)

where:

the coefficient 40.3 is in the unit of m3/s2

TEC :	Total Electron Content along the signal path

c :	speed of light

fd and fu :	downlink and uplink frequencies, respectively.

Example: For a high TEC of 1  1018 electrons/m2 and for fu  14.5 GHz and fd  12.5 GHz this ionospheric delay is equal to 0.859 ns _ 0.639 ns  0.220 ns. So the difference  ½[SPU(k) – SPD(k)] is typically smaller than –0.11 ns.

## 3.5	Tropospheric correction

The troposphere gives a delay depending on the elevation angle, the water content of the air, air density and temperature, but up to 20 GHz this delay is only frequency dependent to a very small extent. So its influence on the difference between the up and down propagation delays is < 10 ps.

## 3.6	Earth station delay measurement

The difference of the transmit and receive section [TX(k) – RX(k)] including the up and down converters, modulator and demodulator (modem), amplifiers, antenna, cabling, etc., has to be determined at each station. Methods to obtain this are:

–	co-location of both stations; or

–	subsequent co-location of a third (transportable) earth station at both stations;

–	the use of a calibrator, e. g., a satellite-simulator and calibrated cable.

Calibrations using a transportable station have found wide applications in the time and frequency community. Here its is assumed that the internal delays of the travelling station are unchanged irrespective of the location were it is installed.

The last method is the least expensive and can be used frequently. This method consists of the calibration of an auxiliary cable, measurement of the sum of the transmit and receive delay, measurement of the sum of the auxiliary cable delay and the receive delay and calculation of the receive and transmit delay from the measurements.

The internal transmit and receive delay difference of the modem has to be determined as well. This can be done by:

–	co-locating the modems and measuring the sum of the transmit delay of one modem and the receive delay of the other;

–	by measuring the sum of transmit and receive delay by connecting the IF output signal to the IF input of each of the modems. The transmit delay between the 1PPS transmitted and the appropriate phase reversal in the IF phase modulated output signal of the modem is measured, e.g. by an oscilloscope. The receive delay is found by subtracting the transmit delay from the measured sum of the delays.

# 4	Combination of TWSTFT measurement data with local measurements

In order to enable the calculation of time-scale differences, each station makes available the TI(k) data together with the assessment of its own corrections, as far as possible, to the other stations. One option is the report of the raw 1-s data collected during a particular session in conjunction with a report of the measured local offsets between the various signals involved (see next paragraph). The other option which has become common practice is to generate a single averaged data point, representative of the session, and only to process this. Independent thereof, one must carefully measure and document the delays which occur in the timing chain of each laboratory. In many cases, the time scales to be compared will ultimately be the local realizations of Coordinated Universal Time (UTC) in laboratory k, called UTC(k). But the method is of course amenable for comparisons of time scales in general, e.g., for the synchronization of the various elements in the ground segment of a Global Navigation Satellite System. Throughout this Recommendation we refer to the situation that the two stations involved each realize its UTC(k).

In general, there is an offset between UTC(k) and the physical signal involved in the measurement process, 1PPSTX(k), representing TS(k). The difference UTC(k)-1PPSTX(k) is in the simplest form reported as REFDELAY(k).

The way UTC(k) is realized differs considerably between institutes. In some laboratories, UTC(k) is a mathematical time scale. All laboratories, however, have a (master) clock (CLOCK(k)) which is a physical representation of this time scale. Each laboratory determines the relationship [UTC(k) – CLOCK(k)], either by measurement or by computation. CLOCK(k) can be used as input 1PPSREF(k) to the modem. There will likely be an offset  [CLOCK(k) – 1PPSREF(k)] caused by the cabling associated with the connection to the modem. The modem generates a 1PPSTX which is related to the transmitted signal which allows the difference [1PPSREF(k) _ 1PPSTX(k)] to be measured.

To calculate the difference between the UTC time scales of the two laboratories from the measured quantities TS(1) and TS(2)  (see Section  2) the following applies:

UTC(1) – UTC(2) 	TS(1) – TS(2)

+ {[UTC(1) – CLOCK(1)]

+ [CLOCK(1) – 1PPSREF(1)]

+ [1PPSREF(1) – 1PPSTX(1)]}

 {[UTC(2) – CLOCK(2)]

 [CLOCK(2) – 1PPSREF(2)]

 [1PPSREF(2) – 1PPSTX(2)]}.

For abbreviation, the terms in curled brackets are designated REFDELAY(k) as introduced before. Deviations of the afore mentioned procedures have been occasionally used, but are not dealt with in this Recommendation.

# 5	Determination of accuracy and stability of the result

The general guidelines of the International Organization for Standardization (ISO) (1993) “Guide to the expression of uncertainty in measurement” should be followed along with the characterization of performance procedures outlined in Recommendation ITU-R TF.538.

# 6	Performance characteristics of the earth station

There are two levels of earth station performance and operation that must be considered for two-way time transfers. The first level concerns meeting the basic requirements of the modem to produce the desired output phase modulation in the pre-determined schedule, the correct transmission frequency, and power level. The second level concerns the performance characteristics of the earth station as required by the regulatory organizations and satellite operating agencies.

Meeting regulatory requirements do not by themselves guarantee satisfactory two-way operation. Only when the required C / N0 and carrier level C is delivered to the modem one will get satisfactory results. Operating parameters are determined from a link budget. From this link budget one can determine the size of antenna required, the transmit power required and the required noise temperature of the receive system. The link budget may be calculated following the direction and examples given in Section  2.3, Chapter 2 and Section  AN2.1, Annex 2 of the ITU Handbook on Satellite Communications (third edition, Geneva, 2002).

Regulatory requirements usually originate with the administrative bodies responsible for the general management of the electromagnetic spectrum at national and international level and with the satellite operator. The rules and regulations often dictate when and where an earth station may be erected, the quality of that earth station, and its ownership and operation. These rules and regulations usually are intended to guarantee that the earth station will not generate interference to other earth stations and to other adjacent satellites. Issues involved usually include antenna patterns and antenna gain, G / T, e.i.r.p., polarization discrimination criteria, and the training and licensing of operating staff. Much of the technical matter concerned with the earth station can be guaranteed by the procurement of systems from manufacturers who have already qualified the equipment for operation with specific satellite systems and standards.

Electrical performance standards are usually defined by one or more of the following: Recommendations ITU-R S.580 and ITU-R S.465; US FCC regulation 25-209 from the Code of Federal Regulations, Title 47, Parts 20-39; and INTELSAT Earth Station Standard documents for standard C (IESS-203), standard E (IESS-205), standard G (IESS-601), and standard Z (IESS-602) earth stations. Some regions of the world may have other requirements relating to the satellite system, location, classification of user, and other criteria.

# 7	TWSTFT operations in a network

The need for standardizing measuring procedures, data processing, and the formats for the exchange of data and relevant information between participating stations is obvious if a network of stations exist and if the data shall be amenable for automatic processing by the involved institutions and third parties. This Recommendation describes well established practice currently in use in networks comprising institutes in Europe and the USA, in Europe and Asia, and also within the Asia-Pacific Region. For each of the networks, a bookkeeper has been assigned, and the work is performed under the auspices of the Working Group on TWSTFT of the Consultative Committee for Time and Frequency (CCTF).

## 7.1	Station identification and assigned characters, codes and offset-frequencies

Participating stations are each identified by a laboratory identification string and an ASCII character. A PRN code and a clean carrier offset frequency are also assigned. The latter is needed when all stations of a network transmit a clean carrier (unmodulated) signal which is part of the operational practice to adjust the receive frequency at the remote station, to monitor the power levels in use, and to allow unambiguous identification of the transmitting stations at a given time. If one institute routinely operates more than one TWSTFT station, each one should get an unique set of identifying parameters.

## 7.2	Session parameters

A plan for a session should contain the following information:

–	satellite identification: name, longitude, orbital parameters, transponder channel, satellite equipment non-reciprocity (if known);

–	station designations;

–	TX-codes and RX-codes;

–	link budget;

–	earth station TX and RX frequencies;

–	point of contact at each station;

–	start date and time;

–	preparation time for each session (equal to the pause between successive sessions);

–	duration of each session;

–	In the report of data, the sequence of collected data points is designated a track, and one speaks of the nominal and actual track length etc..

## 7.3	Local earth station(k) parameters

Each station should keep on file the following information:

- antenna coordinates (x, y, z in the IERS terrestrial reference frame);

–	earth station designator;

–	transmit power e.i.r.p. (W);

- G / T of the receiver (dB/K);

- Nominal C/N0 at the receiver (dBHz);

–	modem: manufacturer, model, type, serial number;

–	time interval counter: manufacturer, type, serial number;

–	antenna: manufacturer, type, diameter, gain;

–	delay calibration: date, method, results;

–	optional: record during the sessions the measured TX-RX delay, received power, the receive C/N0 and frequency, and the environmental parameters at the antenna (temperature, humidity, barometric pressure, weather conditions).

Any change in station parameters and equipment should be reported to the network bookkeeper.

# 8	Calibration of station delays or links

Calibration of the signal delays in the laboratory’s equipment is instrumental for performing “true” time transfer. Without that, assuming stability of the (unknown) delays, one can still perform frequency transfer. Up to now, TWSTFT links have been calibrated by using a portable TWSTFT station, and an uncertainty of  1 ns has been repeatedly demonstrated. Another possibility is the calibration of a TWSTFT link by using an independent time transfer system, such as GPS, or a portable clock. In principle one can distinguish between calibrations of a link – involving two stations -, and of individual signal delays in a particular station.

In order to document the calibration results unambiguously in the TWSTFT data files (to be described in Annex 2), each calibration is characterized by its type, by its calibration identifier CI, by a so-called calibration switch S, and by a calibration result CALR. These paramenter are later referred to in Annex 2.

TYPE	Type of technique used for a certain link characterized by one of the following key words:

PORT ES REL	Portable earth station used in a relative mode.

PORT ES ABS	Portable earth station used in an absolute mode.

PORT SS REL	Portable satellite simulator used in a relative mode.

PORT SS ABS	Portable satellite simulator used in an absolute mode.

Calibration by an independent time transfer system or other means, for example:

GPS		Global Positioning System.

PORT CLOCK	Portable clock.

CIRCULAR T	Monthly bulletin of the BIPM reporting the difference 	between UTC and the local realizations UTC(k).

TRANGLE CLOSURE	Calibration of one link by making use of the calibration results of the links connecting the two stations involved to a third station.CALR	Calibration result in ns. Its use in calcultating times-cale differences is explained in the Section 9 of this Annex 1.

CI	The assignment of a calibration identification is coordinated by the CCTF Working Group on Two-Way Satellite Time and Frequency Transfer for the laboratories involved in a specific calibration campaign. Using an independent time transfer system (e.g. GPS), a calibration is only possible between a pair of laboratories, therefore calibration identifications have to be assigned for pairs of laboratories.

In the data files, the calibration identifier CI is used as a pointer to the file header, which informs about the calibration type, the Modified Julian Day of the first day when data are taken at a respective site in the framework of the calibration, and the estimated uncertainty of the calibration. The data files should only contain the currently relevant calibration entries.

Uncalibrated links should get the indication “999”, and there should be no entry of the link in the header of the file reporting the data (see Annex 2, Section 2.3.3).

If a link is re-established after an interruption, but the knowledge of a previous calibration is carried forward by other means (“bridging”) with a some uncertainty, then a new CI shall be used and TYPE shall designate the old CI and the fact of bridging.

S	The calibration switch (either “0”, “1”, “2”, or “9”) indicates which terms of the two-way equation are included in the calibration result CALR and which equation must therefore be used for the computation of the clock differences (see Section  9.2).

S  0	The calibration result CALR gives the difference between the differential earth station delay (transmit part minus receive part) of the laboratory and the differential earth station delay of the calibrating system (collocation of earth stations, satellite simulator). In consequence, the CALR value is specific for each site.

S  1	The calibration result CALR includes all terms of the two-way equation except the time transfer measurements TW and the reference measurements REFDELAY of the local and remote station, respectively. This is the case using an independent time transfer system, e.g. GPS. In consequence, the CALR value is specific for each link, and in particular CALR(1,2) = -CALR(2,1).

S = 2	The calibration switch S = 2 should be assigned to ranging measurements

S = 9		The calibration switch S = 9 should be used to indicate that no calibration has been performed.

# 9	Calculation of clock differences

## 9.1	Introduction

At present there exist more than one type of modem for TWSTFT, and depending on the modem used and its configuration, TWSTFT data can be reported in two distinct ways.

Data type 1: Each station reports its own measured results, this will be referred to subsequently as “Individual measurements”,

Data type 2: One of the participating station reports in real time the final time transfer results obtained by combining the data measured at both stations. This is referred to as “Combined measurements”

The following subsections explain how clock differences are to be determined based on the reported data from two stations, in dependence on the type of data and on the calibration switch. Reference is made to lines in the data files which are explained in Annex 2. The knowledge of the positions of the earth stations and the satellite allows one to calculate the Sagnac correction (Section 3.2) and the knowledge of the TEC and the transmit and receive frequencies allows one to calculate the  propagation delay due to the ionosphere (Section 3.4).

The reported data can be the raw 1-s measurement data TI(k) collected during a session or the result of a quadratic fit over the data collected during one session, designated TW(k). In the latter case, TW(k) has to be calculated following a strict and unique rule, otherwise the periodic movement of the satellite causes a variation in the time comparison data which is not real. The rule is that the average data point is calculated for the date given by the nominal start date of the session plus half the nominal track length (nominal session duration) rounded to seconds. This rule assumes that the TI(k) data represent instantaneous time interval measurements made at the epoch given in the associated time stamp. If a modem outputs TI(k) as an average over consecutive time interval measurements taken over an interval T (e. g. 1 s), then the epoch of the representative data point TW(k) has to be shifted by T/2. The required information shall be included in the data files (see Annex 2).

## 9.2	Type 1 – Individual measurements

S  0:

UTC(1) – UTC(2) 	0.5[TW(1)  ESDVAR(1)]  REFDELAY(1)	Data line lab 1

–0.5[TW(2)  ESDVAR(2)] - REFDELAY(2)	Data line lab 2

[SCD(2) – SCD(1)]	Header lines lab 1 and lab 2

0.5  [SPU(1)  SPD(1)]	Header line lab 1

+0.5 [SPU(2)  SPD(2)]	Header line lab 2

0.5 [CALR(1)  – CALR(2)]	Data lines lab 1 and lab 2

0.5 XPNDR(1)	Header line lab 1

Here additional quantities have been introduced:

ESDVAR	Monitored differential earth station delay variations (ns) (referred to the differential earth station delay observed at calibration time if a calibration is available). All earth station and modem delay changes have to be included,

XPNDR(k) 	SPT(1) – SPT(2)).

For each TWSTFT network it should be decided whether the ESDVAR is re-set to zero after each delay calibration or not.

S  1:

UTC(1) – UTC(2) 	0.5[TW(1)  ESDVAR(1)]  REFDELAY(1)	Data line lab 1

–0.5[TW(2)  ESDVAR(2)] – REFDELAY(2)	Data line lab 2

+0.5 [CALR(1)  - CALR(2)]		Data lines lab 1 and lab 2

S = 2: Determination of the range to the satellite:

RNG(k) = 0.5 c [TW(k) – CALR(k) – ESDVAR(k)]

Here c has to be understood as the signal propagation speed along the signal path from the satellite to the ground station.

S = 9:

UTC(1) – UTC(2) 	0.5[TW(1)  ESDVAR(1)]  REFDELAY(1)	Data line lab 1

–0.5[TW(2)  ESDVAR(2)] – REFDELAY(2)	Data line lab 2

Note that all above relations remain valid if TW(k) is replaced by TI(k).

## 9.3	Type 2 – Combined measurements

S  0: UTC(LABm) – UTC(LABn) 	ESm  ESn  REFDELAY1

– REFDELAY2

 SAGNAC EFFECT

 IONOSPHERIC CORR.

 CALR

S  1: UTC(LABm) – UTC(LABn) 	ESm  ESn  REFDELAY1

– REFDELAY2

CALR

Annex 2

Description of  the data format for data exchange

# 1	Introduction

The data file begins with a header which should contain all data that are not considered to change during the period of reporting TWSTFT in one file (e. g. one day). The subsequent data lines contain all data that are expected to change from session to session. Some data lines contain information for experimental purposes (e.g. the once-per-second readings) whereas other data lines contain information directly for operational use (e.g. the quadratic least-squares for a session).

Distinction is made between the two types of data as described in Annex 1, Section 9.1. Another distinction is whether the individual 1-s data are reported or the result of a quadratic fit to the 1-s data collected during  the session.

# 2	Data format

## 2.1	General

The conventions and symbols reported below apply completely to the data format described in Section  2.2 and partly to that in Section  2.3 of this Annex.

LAB:	laboratory identification (preferentially according to the Bureau international des poids et mesures (BIPM) convention or as assigned by the network bookkeeper), four characters maximum

jjjjj:	modified Julian day

hh:	UTC hour

mm:	UTC minute

ss:	UTC second

L:	designation of Local laboratory by an ASCII character

R:	designation of Remote laboratory by an ASCII character

*:	indication of the start of a line of text

+:	indication of the sign of the subsequent value, either “+” or blank for a positive value, and “-“ for a negative value

n:	placeholder for a numerical character between 0 and 9

[ ]:	designation of an option

n.nnnnnnnnnnnn:	value of a time interval (s) (12 decimals, corresponding to 1 ps resolution). The type 1 measurement values are typically between 0.250 s and 0.280 s.

## 2.2	Report of the individual 1 s measurements

Only type 1 data are generated of this kind. The data file contains a HEADER followed immediately by DATA lines. For the ease of implementation, rulers have been added in the field descriptions below which are not to become part of the files.

DATA FILE NAME:

Ljjjjjhh.mmR

where jjjjj, hh, mm give the NOMINAL start date of the TWSTFT session.

HEADER:

1         2         3         4         5         6

123456789012345678901234567890123456789012345678901234567890

*    UTC (LAB) – CLOCK	= +n.nnnnnnnnnnnn [jjjjj hhmmss]

*    CLOCK – 1PPSREF	    = +n.nnnnnnnnnnnn [jjjjj hhmmss]

*    1PPSREF – 1PPSTX	= +n.nnnnnnnnnnnn [jjjjj hhmmss]

*    T/2 				    =  +n.nnn [s]

*

*	Any other parameter of the form: PARAMETER = Value [units] [jjjjj hhmmss]

where jjjjj, hhmmss optionally give the date at which the indicated value is taken.

DATA:

1         2         3

123456789012345678901234567890

| jjjjj | hhmmss | n.nnnnnnnnnnnn |
| --- | --- | --- |
| jjjjj | hhmmss | n.nnnnnnnnnnnn |
| jjjjj | hhmmss | n.nnnnnnnnnnnn |

where jjjjj, hhmmss give the date at which the data is taken, and the measurement quantity is TI(k) = [1PPSTX – 1 PPSRX].

EXAMPLE:

File A4926610.56B (data measured at station A during  a TWSTFT session with station B on MJD 49266, scheduled at 10:56 h UTC):

*	A4926610.56B

*	UTC (LAB A) – CLOCK	 – 0.000000123456 49266 101000

*	CLOCK – 1PPSREF	  0.000000012345

*	1PPSREF – 1PPSTX	  0.000000001234 49266 102059

*	DATA  1PPSTX – 1PPSRX (s)

49266	105616 0.270924666406

49266	105617 0.270924663805

49266	105618 0.270924660170

49266	105619 0.270924657628

49266	105620 0.270924654270

49266	105621 0.270924651106

Note that in this example the first data line is generated only at 10:56:16 UTC, which may happen if the modem fails to lock to the received signal during the first seconds of the session.

OP comment: OP considers that the generated output files already done can be used to improve the measurement, as:

1-) ITU file named TW…. will be kept as it is (not to modify)

2-) 1s data file can be modified by adding three additional columns (one for the IF lock, one for the received power, and one for the carrier to noise at the receiver). All the three parameters can be provided by the modem. The name of the 1s data file will not be changed but should be available to the users every day. All the 1s data files for a day can be compressed in order to generate only one file per day and per station, and be named as T1SLjjjjj.ZIP (T for two-way, 1S for 1s data, L for local station, jjjjj for MJD). This zip file will contain all the 1s data files for one day, for each session, and for each remote lab.

To resume: each lab will generate only two files per day: as an example for OP at MJD 49266, the two generated files will be: TWOP49.266 and T1SOP49266.ZIP. The zip file will contain all the files A49266hh.mmR (R for remote station). It remains to each lab, after uploading the files from remote stations, to extract the needed files for any remote station.

INRIM, NIST and PTB are currently not in favour of this option

## 2.3	File format reporting results of a quadratic fit

Examples in Section 3 of this Annex 2 illustrate the format described subsequently.

### 2.3.1	General remarks

The purpose of this format is to reduce the amount of data to be exchanged and to be able to report in one data file tracks of one laboratory involving different partner stations and different satellite links. Data of more than one day may be reported in one file. It allows clock differences to be calculated in an easy way, using the information given in the header and data lines, without having to know individual measurement set-ups at the participating laboratories.

There are two kinds of lines:

–	lines with an asterisk in column one (file header, data line header);

–	lines without an asterisk in column one (data lines).

In the format description, characters in bold letters are keywords at a certain position, characters in cursive letters have to be replaced by actual strings or values, respectively. Strings in brackets are optional, and whenever data must be preceded by a sign, it is indicated by a “”. Any missing data should be replaced by series of “9” covering the full length of the respective data field, including the field reserved for the sign and the decimal point.

### 2.3.2	File name

#### 2.3.2.1	Type 1 – Individual measurements

The file name consists of TW, the designation of the laboratory (LLLL, laboratory identification, preferentially according to the Bureau international des poids et mesures (BIPM) convention or as assigned by the network bookkeeper, up to four characters)  and the Modified Julian Day as given in the first data line (the last three digits as file extension): TWLLLLMM.MMM (for example TWTUG50.091). The same rule shall apply of data from more than one day are reported in the file. In case that two ore more Earth stations are operated at one laboratory, the data can be reported in one file, as each data line unambiguously refers to the stations in use. However, it may be more convenient to report an individual file for each station. In this case the LLLL in the file name shall allow distinction between the stations used, e.g., PTB1 and PTB2, or USNO and USNP, etc., as agreed with the network bookkeeper.

#### 2.3.2.2	Type 2 – Combined measurements

The file name consists of TW (for two-way satellite time and frequency transfer), the designation of the region of the laboratory pair (LLLL, up to four characters) and the Modified Julian Day as given in the first data line (the last three digits as file extension): TWLLLLMM.MMM (for example TWAP52.520).

### 2.3.3	Header

Number of columns in header lines are limited to 78.

.........1.........2.........3.........4.........5.........6.........7........
123456789012345678901234567890123456789012345678901234567890123456789012345678

* TWLLLLMM.MMM
* FORMAT	nn
* LAB	LLLL
* REV DATE	YYYY-MM-DD
* ES LLLLnn	LA: D dd mm ss.sss   	LO: D ddd mm ss.sss	HT: +nnnn.nn m
* REF-FRAME	RRRRRRRRRR
* LINK  LL	SAT: SSSSSSSSSSSSSSSSSS	NLO: D ddd mm ss.sss	XPNDR: +nnnn.nnn ns

*	SAT-NTX: fffff.ffff MHz	SAT-NRX: fffff.ffff MHz

* CAL  CCC	TYPE: TTTTTTTTTTTTTTTTT	MJD: MMMMM  EST. UNCERT.: nnnn.nnn ns
* LOC-MON	[YES] [NO]
* MODEM	TYPE, SERIAL NUMBER
* COMMENTS	SSSSSSSSSSSSSSSSSSSSSSSSSSSSSSSSSSSSSSSSSSSSSSSSSSSSSSSSSSSSSSSSSS

*

TWLLLLMM.MMM	File name (see Section  2.3.2)

FORMAT			Format version (from 01 to 99).

LAB				Laboratory identification, preferentially according to the Bureau international des poids et mesures (BIPM) convention or as assigned by the network bookkeeper.

REV DATE			Revision date of the header data, changed when one parameter given in the header is changed apart from the file name. Each file of one laboratory has to have a unique file name and different files may have the same header. YYYY-MM-DD for year, month and day.

ES				Name and position of earth station:

LLLLnn	Name of earth station formed of the laboratory's acronym (up to four characters) and a station identification number (two digits), for example TUG01. See Section  2.3.2.1 for the naming of the data file.

LA	Latitude in geodetic coordinates, replacing D by N (North) or S (South).

LO	Longitude in geodetic coordinates, replacing D by W (West) or E (East).

HT	Height (m).

REF-FRAME	Designation of the reference frame of the two-way antenna coordinates. As many columns as necessary:

LINK	LINK lines characterize the satellite links.

LL	Link identification contained in each data line (data line header: LI, see Section  2.3.4), which points to a specific LINK line of the file header.

SAT	Satellite identification according to the designation of the satellite operating agency.

NLO	Nominal longitude of satellite, replacing D by W (West) or E (East).

XPNDR	Differential transponder delay (ns) (local station to remote station minus remote station to local station).

SAT-NTX	Nominal transmit frequency of the satellite (MHz). This frequency corresponds to the receive frequency of the local earth station.

SAT-NRX	Nominal receive frequency of the satellite (MHz). This frequency corresponds to the transmit frequency of the local earth station.

Two lines describe one link, but as many couples of LINK lines as necessary can be used.

CAL	CAL lines give the calibration type, the Modified Julian Day and the estimated uncertainty of the calibration:

CCC	Calibration identification contained in each data line (data line header: CI, see Section  2.3.4), which points to a specific CAL line of the file header.

TYPE	Type of technique used for a certain link characterized by one of the following key words as described in Annex 1 Section  9.

Different types may be used for the same link resulting in different data lines.

MJD			Modified Julian Day of the calibration.

EST.UNCERT			Estimated uncertainty of the calibration.

A calibration is described in one line, but as many CAL lines as necessary can be used.

LOC-MON	YES or NO has to be used depending on the availability of a local earth station delay monitoring system.

MODEM	Type and serial number of modem used. All modems used during the period of validity of the file should be reported.

As many lines and columns (up to 78 columns, see 2.3.3) as necessary can be used.

COMMENTS	Comment lines as many as necessary.

The last line of the file header contains an asterisk at column 1 only.

### 2.3.4	Data line

#### 2.3.4.1	Type 1 – Individual measurements

For the format of a data line see Example 1 at the end of this Annex..

EARTH-STAT	Designation of local (LOC) and remote (REM) earth station given by the respective laboratory’s acronym (up to four characters) and a station identification number (two digits), for example TUG01. This field is right justified.

LI	Link identification is a pointer to the file header that gives information on a specific satellite link.

The assignment of link identifications has to be coordinated on an international basis such as by the CCTF Working Group on Two-Way Satellite Time and Frequency Transfer .

MJD STTIME	Nominal start date (Modified Julian Day and time (hour, minute, second), referenced to UTC).

NTL	Nominal track length (s) (nominal date of last sample minus nominal date of first sample).

TW	Result (s), of a quadratic fit over the data of one track calculated for the date given by the nominal start date plus half of the nominal track length rounded to seconds (corresponds to TW(k) from Annex 1, Section 9.2).

DRMS	Root mean square of the residuals to the quadratic fit (ns).

SMP	Number of samples used in the quadratic fit.

ATL	Actual track length (s) (date of last sample minus date of first sample of the quadratic fit).

REFDELAY	Delay of the reference with respect to UTC(LAB) (s), explained in Annex 1, Section 4.

RSIG	Uncertainty of the REFDELAY (ns).

CI	Calibration identification is a pointer to the file header.

S	Calibration switch (either “0”, “1”, “2” or “9”).

CALR	Calibration result (ns).

ESDVAR	Monitored differential earth station delay variations (ns).

ESIG	Uncertainty of ESDVAR (ns).

TMP	Outside temperature (C).

HUM	Outside relative humidity (%).

PRES	Air pressure (hPa).

#### 2.3.4.2	Type 2 – Combined measurements

For the format of a data line see Example 5 in Section 4of this Annex.

EARTH-STAT	Designation of local (LOC) and remote (REM) earth station given by the respective laboratory’s acronym (up to four characters) and a station identification number (two digits), for example CRL01.

LI	Link identification is a pointer to the file header that gives information on a specific satellite link.

The assignment of link identifications has to be coordinated on an international basis such as by the CCTF Working Group on Two-Way Satellite Time and Frequency Transfer .

MJD STTIME	Nominal start date (MJD and time (h, min, s) referenced to UTC).

NTL	Nominal track length (s) (nominal date of last sample minus nominal date of first sample).

ESm-ESn	Result (s), of a quadratic fit and inner delay of modem (ns).

DRMS	Root mean square of the residuals to the quadratic fit (ns).

SAGNAC	Sagnac effect (ns).

UTCm-ESm	Delay from UTC(m) reference point to 1PPS input point of modem (ns).

UTCn-ESn	Delay from UTC(n) reference point to 1PPS input point of modem (ns).

CI	Calibration identification

S	Calibration switch (either “0” or “1”).

CALR	Calibration result (ns).

TMPm	Outside temperature (C) at m-station.

HUMm	Outside relative humidity (%) at m-station.

PRESm	Air pressure (hPa) at m-station.

TMPn	Outside temperature (C) at n-station.

HUMn	Outside relative humidity (%) at n-station.

PRESn	Air pressure (hPa) at n-station.

# 3	Examples

The examples contain actual and fictitious data.

3. Examples Example 1	Data line for reporting individual data, quadratic fit.

0         0         0         0         0         0         0         0         0         1         1         1         1

1         2         3         4         5         6         7         8         9         0         1         2         3

1234567890123456789012345678901234567890123456789012345678901234567890123456789012345678901234567890123456789012345678901234567890

* EARTH-STAT  LI  MJD  STTIME NTL        TW        DRMS SMP ATL      REFDELAY    RSIG  CI S    CALR     ESDVAR   ESIG TMP HUM PRES

* LOC    REM           hhmmss  s         s          ns       s          s         ns            ns        ns      ns degC  %  mbar

LLLLnn LLLLnn LL MMMMM hhmmss nnn n.nnnnnnnnnnnn n.nnn nnn nnn n.nnnnnnnnnnnn n.nnn CCC i nnnn.nnn nnnn.nnn n.nnn nn nnn nnnn

# Example 2 and 3:	Sample data files from a station in Europe (2) and in USA (3)

2: TWSTFT file of PTB, MJD 54710, 2008-09-01

* twptb54.710

* FORMAT    01

* LAB       PTB

* REV DATE  2008-08-28

* ES  PTB01 LA: N  52 17 49.787      LO: E  10 27 37.966   HT:   143.41 m

* REF-FRAME WGS84

* LINK   07 SAT: INTELSAT 707        NLO: E 307 00 00.000  XPNDR:     0.000 ns

*           SAT-NTX: 12518.2000 MHz  SAT-NRX: 14013.2000 MHz

* LINK   09 SAT: INTELSAT 707        NLO: E 307 00 00.000  XPNDR: +9999.999 ns

*           SAT-NTX: 12716.7500 MHz  SAT-NRX: 14211.7500 MHz

* LINK   10 SAT: INTELSAT 3R         NLO: E 317 00 00.000  XPNDR:     0.000 ns

*           SAT-NTX: 12574.2500 MHz  SAT-NRX: 14072.2500 MHz

* LINK   11 SAT: INTELSAT 3R         NLO: E 317 00 00.000  XPNDR: +9999.999 ns

*           SAT-NTX: 12627.0500 MHz  SAT-NRX: 14330.7500 MHz

* CAL   113 TYPE: CIRCULAR T         MJD: 54525  EST. UNCERT.:    5.200 ns

* CAL   114 TYPE: CAL 083 BRIDGED    MJD: 54502  EST. UNCERT.:    2.000 ns

* CAL   115 TYPE: CAL 103 BRIDGED    MJD: 54502  EST. UNCERT.:    1.300 ns

* CAL   116 TYPE: CAL 096 BRIDGED    MJD: 54502  EST. UNCERT.:    1.100 ns

* CAL   117 TYPE: CAL 109 BRIDGED    MJD: 54502  EST. UNCERT.:    1.100 ns

* CAL   118 TYPE: CIRCULAR T         MJD: 54502  EST. UNCERT.:    5.000 ns

* CAL   119 TYPE: CAL 106 BRIDGED    MJD: 54502  EST. UNCERT.:    1.200 ns

* CAL   120 TYPE: CAL 107 BRIDGED    MJD: 54502  EST. UNCERT.:    1.200 ns

* LOC-MON   NO

* MODEM     SATRE 037

* COMMENTS

*

* EARTH-STAT  LI  MJD  STTIME NTL        TW        DRMS SMP ATL     REFDELAY     RSIG  CI S    CALR     ESDVAR   ESIG TMP HUM PRES

* LOC    REM           hhmmss  s         s          ns       s         s          ns            ns        ns      ns degC  %  mbar

PTB04  PTB04 10 54710 000700 119  0.268701755755 0.375 120 119  0.000001981575 0.009 999 9 999999999    -0.180 0.100  18  61 1002

PTB04   IT02 10 54710 001300 119  0.266832337354 0.607 120 119  0.000001981520 0.010 116 1   316.100    -0.180 0.100  18  61 1002

PTB04  ROA01 10 54710 001600 119  0.262320415926 0.448 120 119  0.000001981613 0.016 118 1   288.400    -0.180 0.100  17  62 1002

PTB04   OP01 10 54710 001900 119  0.266437968645 0.376 120 119  0.000001981576 0.014 115 1  7316.500    -0.180 0.100  17  62 1002

PTB04  NPL01 10 54710 002200 119  0.266586507349 0.693 120 119  0.000001981630 0.014 999 9 999999999    -0.180 0.100  17  62 1002

PTB04   CH01 10 54710 003700 119  0.267009789103 0.375 120 119  0.000001981722 0.018 117 1   204.550    -0.180 0.100  17  64 1001

PTB04  IPQ01 10 54710 004000 119  0.262213386716 0.529 120 119  0.000001981839 0.012 999 9 999999999    -0.180 0.100  17  64 1001

PTB04  AOS01 10 54710 004300 119  0.270015381981 0.415 120 119  0.000001981716 0.012 999 9 999999999    -0.180 0.100  17  64 1001

PTB04 USNO01 11 54710 004600 119  0.262319009268 0.224 120 119  0.000001981668 0.009 114 1  -218.800    -0.180 0.100  17  65 1002

PTB04 NIST01 11 54710 004900 119  0.268893360924 0.225 120 119  0.000001981639 0.013 113 1    30.100    -0.180 0.100  17  65 1002

3:TWSTFT file of NIST, MJD 54701 (2008-08-23)

* TWNIST54.700

* FORMAT    01

* LAB       NIST

* REV DATE  2008-08-22

* ES NIST01 LA: N  39 59 45.000      LO: W 105 15 46.000   HT: +1640.00 m

* REF-FRAME WGS84

* LINK   11 SAT: INTELSAT 3R         NLO: E 317 00 00.000  XPNDR: 999999999 ns

*           SAT-NTX: 12030.7500 MHz  SAT-NRX: 14375.0500 MHz

* CAL   071 TYPE: Cir.T(NPL1)        MJD: 53152  EST. UNCERT.:    5.000 ns

* CAL   072 TYPE: Cir.T(IEN/IT)      MJD: 52989  EST. UNCERT.:    5.000 ns

* CAL   063 TYPE: Cir.T(PTB)         MJD: 52239  EST. UNCERT.:    5.000 ns

* CAL   070 TYPE: Cir.T(PTB)         MJD: 52869  EST. UNCERT.:    5.000 ns

* CAL   007 TYPE: Cir.T(TUG)         MJD: 50599  EST. UNCERT.:    5.000 ns

* CAL   008 TYPE: Cir.T(VSL)         MJD: 51389  EST. UNCERT.:    5.000 ns

* CAL   085 TYPE: Cir.T(PTB)         MJD: 53643  EST. UNCERT.:    5.000 ns

* CAL   113 TYPE: Cir.T(PTB)         MJD: 54525  EST. UNCERT.:    5.200 ns

* CAL   322 TYPE: TRIANGLE CLOSURE   MJD: 54584  EST. UNCERT.:    6.000 ns

* CAL   324 TYPE: TRIANGLE CLOSURE   MJD: 54584  EST. UNCERT.:    6.000 ns

* CAL   326 TYPE: TRIANGLE CLOSURE   MJD: 54584  EST. UNCERT.:    6.000 ns

* CAL   328 TYPE: TRIANGLE CLOSURE   MJD: 54584  EST. UNCERT.:    6.000 ns

* CAL   329 TYPE: TRIANGLE CLOSURE   MJD: 54584  EST. UNCERT.:    6.000 ns

* CAL   330 TYPE: TRIANGLE CLOSURE   MJD: 54584  EST. UNCERT.:    6.000 ns

* CAL   331 TYPE: TRIANGLE CLOSURE   MJD: 54584  EST. UNCERT.:    6.000 ns

* CAL   999 TYPE: UNCALIBRATED       MJD: 99999  EST. UNCERT.: 99999999 ns

* LOC-MON   NO

* MODEM     SATRE, S/N 78

*

* EARTH-STAT  LI  MJD  STTIME NTL        TW        DRMS SMP ATL     REFDELAY     RSIG  CI S    CALR     ESDVAR   ESIG TMP HUM PRES

* LOC    REM           hhmmss  s         s          ns       s         s          ns            ns        ns      ns degC  %  mbar

NIST01  IPQ01 11 54700 001900 119 +0.262387824848 0.246 120 119 +0.000000860500 99999 999 9 999999999 00224.040 0.200  35   1  829

NIST01   CH01 11 54700 002800 119 +0.267188689396 0.088 120 119 +0.000000860500 99999 324 1   176.060 00224.040 99999  35   1  829

NIST01   OP01 11 54700 003700 119 +0.266621475071 0.475 120 119 +0.000000860500 99999 328 1  7287.687 00224.040 99999  35   1  829

NIST01  VSL01 11 54700 004300 119 +0.267702142625 0.420 120 119 +0.000000860500 99999 331 1   273.323 00224.040 99999  35   1  829

NIST01  PTB04 11 54700 004900 119 +0.268889783456 0.158 120 119 +0.000000860500 99999 113 1   -30.100 00224.040 99999  35   1  829

NIST01   IT02 11 54700 005200 119 +0.267019344143 0.144 120 119 +0.000000860500 99999 326 1   285.833 00224.040 99999  35   1  829

NIST01  ROA01 11 54700 005500 119 +0.262506935786 0.097 120 119 +0.000000860500 99999 999 9 999999999 00224.040 0.200  35   1  829

# Example 4	Data line for reporting combinded data

0         0         0         0         0         0         0         0         0         1         1         1        1

1         2         3         4         5         6         7         8         9         0         1         2        3

1234567890123456789012345678901234567890123456789012345678901234567890123456789012345678901234567890123456789012345678901234567890

* EARTH-STAT  LI  MJD  STTIME  NTL  ESm-ESn   DRMS SAGNAC  UTCm-ESm  UTCn-ESn CI  S    CALR   TMPm  HUMm  PRESm   TMPn  HUMn  PRESn

*  LOC   REM           hhmmss   s     ns      ns    ns       ns        ns             ns     degC   %     hPa    degC   %     hPa

LLLLnn LLLLnn LL MMMMM hhmmss nnn +nnnn.nn  n.nnn +nn.nn  +nnnn.nn +nnnn.nn  CCC i +nnnn.nnn +nn   nnn   nnnn    +nn   nnn   nnnn

# Example 5:	Sample data files from a station in Asia, reporting combined data

TWAP53.440 from NICT, MJD 53440, 2005-03-11

* FORMAT    99

* LAB       NICT

* REV DATE  2005-03-11

* ES1 NICT01LA: N  35 42 28.194      LO: E 139 29 16.527   HT:    93.76 m

* REF-FRAME GRS80

* ES7 NICT03LA: N  35 42 28.194      LO: E 139 29 16.527   HT:    93.76 m

* REF-FRAME GRS80

* ES3 AUS01 LA: S  33 46 58.258      LO: E 151  9  5.945   HT:   101.44 m

* REF-FRAME WGS84

* ES4 NMIJ01LA: N  36  4  0.000      LO: E 140  8  0.000   HT:    90.00 m

* REF-FRAME WGS84

* ES5 NTSC01LA: N  34 22 12.000      LO: E 109 13 12.000   HT:   467.00 m

* REF-FRAME WGS84

* ES6 TL01  LA: N  24 57 13.000      LO: E 121  9 52.000   HT:   203.00 m

* REF-FRAME WGS84

* LINK   12 SAT: PAS-8               NLO: E 166  0  0.000  XPNDR: 999999999 ns

*           SAT-NTX1:12417.0000 MHz  SAT-NRX1:14226.0000 MHz  (for NICT)

*           SAT-NTX2:12476.0000 MHz  SAT-NRX2:14167.0000 MHz  (for AUS)

* CALR  999 TYPE: 99999999999999999  MJD: 99999  EST. UNCERT.: 9999.999 ns

* LOC-MON1  NO

* LOC-MON2  NO

* MODEM     ATLANTIS

*

* LINK   10 SAT: JCSAT-1B            NLO: E 150  0  0.000  XPNDR: 999999999 ns

*           SAT-NTX: 12306.0000 MHz  SAT-NRX: 14054.0000 MHz

* CALR  202 TYPE: CIRCULAR T         MJD: 52562  EST. UNCERT.:   10.000 ns

* LOC-MON1  NO

* LOC-MON2  NO

* MODEM     ATLANTIS

*

* LINK   10 SAT: JCSAT-1B            NLO: E 150  0  0.000  XPNDR: 999999999 ns

*           SAT-NTX: 12306.0000 MHz  SAT-NRX: 14054.0000 MHz

* CALR  999 TYPE: 99999999999999999  MJD: 99999  EST. UNCERT.: 9999.999 ns

* LOC-MON1  NO

* LOC-MON2  NO

* MODEM     ATLANTIS

*

* LINK   10 SAT: JCSAT-1B            NLO: E 150  0  0.000  XPNDR: 999999999 ns

*           SAT-NTX: 12306.0000 MHz  SAT-NRX: 14054.0000 MHz

* CALR  204 TYPE: CIRCULAR T         MJD: 52194  EST. UNCERT.:   10.000 ns

* LOC-MON1  NO

* LOC-MON2  NO

* MODEM     ATLANTIS

*

* EARTH-STAT  LI  MJD  STTIME NTL  ESm-ESn  DRMS  SAGNAC   UTCm-ESm  UTCn-ESn  CI S      CALR  TMPm HUMm PRESm  TMPn HUMn PRESn

*  LOC    REM          hhmmss   s       ns    ns      ns         ns        ns              ns  degC    %   hPa  degC    %   hPa

NICT03  AUS01 12 53440 005100 295 -2096.70  0.23   33.24    1395.79     78.45 999 1  9999.999  99.9  999  9999  99.9  999  9999

NICT03  AUS01 12 53440 005800 294 -2096.28  0.22   33.24    1395.79     78.45 999 1  9999.999  99.9  999  9999  99.9  999  9999

NICT01 NMIJ01 10 53440 012200 295 -1712.04  0.15    2.12    1395.79     41.40 202 1   343.070  99.9  999  9999  99.9  999  9999

NICT01 NMIJ01 10 53440 013000 294 -1711.89  0.15    2.12    1395.79     41.40 202 1   343.070  99.9  999  9999  99.9  999  9999

NICT01 NMIJ01 10 53440 013700 294 -1711.55  0.17    2.12    1395.79     41.40 202 1   343.070  99.9  999  9999  99.9  999  9999

NICT01 NMIJ01 10 53440 014500 294 -1711.26  0.16    2.12    1395.79     41.40 202 1   343.070  99.9  999  9999  99.9  999  9999

NICT01 NTSC01 10 53440 020100 295 -1500.47  0.26  -85.29    1395.79    180.73 999 1  9999.999  99.9  999  9999  99.9  999  9999

NICT01 NTSC01 10 53440 020900 294 -1500.22  0.25  -85.29    1395.79    180.73 999 1  9999.999  99.9  999  9999  99.9  999  9999

NICT01 NTSC01 10 53440 021600 294 -1500.49  0.28  -85.29    1395.79    180.73 999 1  9999.999  99.9  999  9999  99.9  999  9999

NICT01   TL01 10 53440 023300 295 -1632.97  0.17  -63.07    1395.79      4.10 204 1   285.632  99.9  999  9999  99.9  999  9999

NICT01   TL01 10 53440 024300 294 -1632.89  0.17  -63.07    1395.79      4.10 204 1   285.632  99.9  999  9999  99.9  999  9999

NICT01   TL01 10 53440 024900 294 -1633.12  0.19  -63.07    1395.79      4.10 204 1   285.632  99.9  999  9999  99.9  999  9999