Physikalisch-Technische Bundesanstalt, Germany 			August 2005

Report to the CCL and to the CCL/CCTF joint working group

on recent frequency measurements

of the 435 nm 2S1/2(F=0) - 2D3/2(F=2) transition of 171Yb+

(i) General conditions of measurement

In an extension of previous work [1-3], four frequency measurements of the 688 THz (435 nm) transition of 171Yb+ were performed in July and August 2005. The output of a frequency-doubled diode laser was locked to the 2S1/2(F=0) - 2D3/2(F=2) transition of a single 171Yb+ ion confined in a Paul trap. The diode laser frequency was compared to the caesium fountain clock CSF1 of PTB using a frequency comb generator based on a Er3+-doped fiber laser [4].

Immediately before one of the absolute frequency measurement intervals, two 171Yb+ traps were operated simultaneously as described in [5] in order to compare the transition frequencies of the trapped ions for various orientations of the magnetic field. The results of this comparison confirm the results reported in [5].

The 171Yb+ trap was operated at room temperature (T=297 K). The 171Yb+ transition frequency values reported here include the frequency shift due to the ambient blackbody radiation. Tabulated oscillator strength data and the atomic polarizability measurements described in [5,6] indicate that the blackbody AC Stark shift of the Yb+ reference transition is -0.37(5) Hz at T=300 K.

(ii) Measurement results and uncertainty contributions

Notation: i(Yb+) = 688 358 979 309 000 Hz + xi Hz,  i: number of measurement

i            Date                    xi / Hz              uAi / Hz        uB(Cs) / Hz     uB(Yb+) / Hz

1         05.07.05                307.84               3.43                1.82                1.05

2         06.07.05                307.51               0.46                1.82                1.05

3         09.08.05                307.49               1.01                1.82                1.05

4         10.08.05                307.07               0.64                1.82                1.05

Weighted mean of measured 171Yb+ 2S1/2(F=0) - 2D3/2(F=2) transition frequencies:

(Yb+) = 688 358 979 309 307.65 Hz

(Weighting proportional to (uAi2 + uBi2(Cs) + uBi2(Yb+))-1.

Earlier results [2,3] are included with <0.3 Hz effect on the mean.)

Type A uncertainty of (Yb+), including earlier results:

uA = (uAi-2)-1/2 = 0.40 Hz

Type B uncertainty of (Yb+), recent measurements:

uB = (uB2(Cs) + uB2(Yb+))1/2 = 2.10 Hz

Combined uncertainty of (Yb+):

u(combined) = (uA2 + uB2)1/2 = 2.14 Hz

(iii) Comments on some type B uncertainty contributions

(see also GUM worksheet, Appendix 1)

(1) The value uB(Cs) = 1.8 Hz (corresponding to a fractional uncertainty of 2.710-15) reflects the present state of a new systematic uncertainty evaluation of CSF1 after modifications of the magnetic field geometry.

(2) The assumed quadrupole shift contribution to uB(Yb+) is a factor of two larger than the statistical uncertainty of the trap-trap-comparison measurements described in [4]. These comparisons did not show any statistically significant quadrupole or tensorial Stark shifts.

(3) The uB(Yb+) contributions taking into account the trap field-induced Stark shift and the relativistic Doppler shift correspond to a stray-field induced micromotion amplitude that is a factor of two larger than the maximum amplitude which might remain undetected with the employed compensation scheme for electrostatic stray fields [5,6]. The amplitude of the thermal secular motion is much smaller than the assumed micromotion amplitude.

(4) The blackbody shift contribution to uB(Yb+) is an estimate of possible deviations of the AC Stark shift from the purely scalar shift caused by isotropic 300 K blackbody radiation. Such deviations could be caused by spurious thermal radiation sources near the trap and by laser stray light. Tests of the mechanical shutters which block the cooling and repumping laser light during the clock laser pulse were carried out during the trap-trap comparison experiments (separate shutters are used for the two trap setups).

(5) The uB(Yb+) contribution taking into account the quadratic Zeeman shift caused by the rf trap drive current is based on an estimate of the displacement current through one octant of the trap electrode structure and of the resulting magnetic field. In a perfectly symmetric trap, the magnetic field contributions of all octants would add up to zero at trap center. The field produced by other rf conductors in the setup is expected to be smaller than the assumed field since these conductors are located sufficiently far from trap center.

References

[1]  J. Stenger, Chr. Tamm, N. Haverkamp, S. Weyers, H.R. Telle,

Opt. Lett. 26, 1589 (2001).

[2]  T. Quinn, Metrologia 40, 103 (2003)

[3]  E. Peik, B. Lipphardt, H. Schnatz, T. Schneider, Chr. Tamm,

Phys. Rev. Lett. 93, 170801 (2004).

[4]  Ph. Kubina, P. Adel, G. Grosche, Th. W. Hänsch, R. Holzwarth, B. Lipphardt, 
      H. Schnatz, Opt. Express 904 (2005)

[5]  T. Schneider, E. Peik, Chr. Tamm, Phys. Rev. Lett. 94, 230801 (2005).

[6]  T. Schneider, PhD thesis, Univ. Hannover (2005).

Appendix 1:

GUM worksheet:

Yb+ frequency standard, Jul/Aug. 2005 measurements,

corrections and systematic uncertainty contributions

Model Equation:

fYbCorr = f2ndZeem + fgrav + QShift + Servo + StarkBBDev + StarkTrap + relDopp;

f2ndZeem = fZeemDC + fZeemAC;

fZeemDC = s0*(BDC)^2;

fZeemAC = (s0/2)*(BAC)^2;

fgrav = s3*hrefminusYb

List of Quantities:

| Quantity | Unit | Definition |
| --- | --- | --- |
| fYbCorr | Hz | corrections to fYb due to interactions of trapped ion |
| f2ndZeem | Hz | time average of second-order Zeeman shift |
| fgrav | Hz | gravitational shift of measured frequency |
| QShift | Hz | uncertainty due to stray-field induced quadrupole shift of Yb frequency |
| Servo | Hz | uncertainty due to servo errors and spectral asymmetry of probe laser |
| StarkBBDev | Hz | uncertainty due to deviation of the AC Stark shift from 300 K-blackbody shift; includes shift due to laser stray light |
| StarkTrap | Hz | uncertainty due to quadratic Stark shift caused by secular motion and by excessive micromotion |
| relDopp | Hz | uncertainty due to relativistic Doppler shift caused by secular motion and by excessive micromotion |
| fZeemDC | Hz | second-order Zeeman shift due to applied static magnetic field |
| fZeemAC | Hz | time average of second-order Zeeeman shift due to magnetic field associated with rf trap drive |
| s0 | Hz/T2 | sensitivity coefficient of quadratic Zeeman shift |
| BDC | T | static magnetic field at trap center |
| BAC | T | amplitude of rf magnetic field at trap center |
| s3 | Hz/m | gravitational shift coefficient of measured frequency |
| hrefminusYb | m | elevation difference of Cs reference and Yb+ standard |

fYbCorr: Result

total systematic shift of Yb+ single-ion standard, August 2005

f2ndZeem: Interim Result

fgrav: Interim Result

gravitational redshift difference due to different elevation of Cs clock and Yb+ standard

QShift: Stray-field induced quadrupole shift

Type B normal distribution

Value: 0 Hz

Expanded Uncertainty: 1 Hz

Coverage Factor: 1

Servo: Servo error

Type B normal distribution

Value: 0 Hz

Expanded Uncertainty: 0.1 Hz

Coverage Factor: 1

StarkBBDev: AC Stark shift minus 300 K blackbody shift

Type B normal distribution

Value: 0 Hz

Expanded Uncertainty: 0.3 Hz

Coverage Factor: 1

StarkTrap: Stark shift due to trap field

Type B rectangular distribution

Value: 0 Hz

Halfwidth of Limits: 0.03 Hz

relDopp: Relativistic Doppler shift due to micromotion and secular motion

Type B rectangular distribution

Value: -0.01 Hz

Halfwidth of Limits: 0.01 Hz

fZeemDC: Interim Result

fZeemAC: Interim Result

s0: Coefficient for quadratic Zeeman shift

Type B normal distribution

Value: 520·108 Hz/T2

Expanded Uncertainty: 26·108 Hz/T2

Coverage Factor: 1

BDC: Applied static magnetic field

Type B normal distribution

Value: 3.09·10-6 T

Expanded Uncertainty: 0.1·10-6 T

Coverage Factor: 1

BAC: AC magnetic field amplitude

Type B rectangular distribution

Value: 2·10-8 T

Halfwidth of Limits: 2·10-8 T

s3: Proportionality factor for gravitational redshift

Type B rectangular distribution

Value: 0.0688 Hz/m

Halfwidth of Limits: 1·10-6 Hz/m

hrefminusYb: Height Fountain- height Yb+

Type B rectangular distribution

Value: -0.75 m

Halfwidth of Limits: 0.01 m

Uncertainty Budget:

| Quantity | Value | Standard Uncertainty | Degrees of Freedom | Sensitivity Coefficient | Uncertainty Contribution | Index |
| --- | --- | --- | --- | --- | --- | --- |
| f2ndZeem | 0.4965 Hz | 0.0406 Hz |  |  |  |  |
| fgrav | -0.051600 Hz | 397·10-6 Hz |  |  |  |  |
| QShift | 0.0 Hz | 1.00 Hz | 50 | 1.0 | 1.0 Hz | 90.7 % |
| Servo | 0.0 Hz | 0.100 Hz | 50 | 1.0 | 0.10 Hz | 0.9 % |
| StarkBBDev | 0.0 Hz | 0.300 Hz | 50 | 1.0 | 0.30 Hz | 8.2 % |
| StarkTrap | 0.0 Hz | 0.0173 Hz |  | 1.0 | 0.017 Hz | 0.0 % |
| relDopp | -0.01000 Hz | 5.77·10-3 Hz |  | 1.0 | 5.8·10-3 Hz | 0.0 % |
| fZeemDC | 0.4965 Hz | 0.0406 Hz |  |  |  |  |
| fZeemAC | 10.4·10-6 Hz | 12.0·10-6 Hz |  |  |  |  |
| s0 | 52.00·109 Hz/T2 | 2.60·109 Hz/T2 | 50 | 9.5·10-12 | 0.025 Hz | 0.1 % |
| BDC | 3.090·10-6 T | 100·10-9 T | 50 | 320·103 | 0.032 Hz | 0.1 % |
| BAC | 20.0·10-9 T | 11.5·10-9 T |  | not valid! | 12·10-6 Hz | 0.0 % |
| s3 | 0.068800000 Hz/m | 577·10-9 Hz/m |  | -0.75 | -430·10-9 Hz | 0.0 % |
| hrefminusYb | -0.75000 m | 5.77·10-3 m |  | 0.069 | 400·10-6 Hz | 0.0 % |
| fYbCorr | 0.4 Hz | 1.05 Hz | 60 |  |  |  |

Result:

Quantity: fYbCorr

Value: 0.4 Hz

Expanded Uncertainty: ±2.1 Hz

Coverage Factor: 2.0

Coverage: t-table 95%